Vnitřek otřít lihem.

1. Otočit klíčkem aby svítil (zůstává v poloze tak, že svítí)
2. Z boku zmáčknout tlačítko, přidržet několik sekund (ukáže se baterka).
3. Znova zmáčknout nabíhá program.
4. Zavřít box až dolu.
5. Zmačknout UV (je nastavená na 20 minut).
6. Otevřít dveře do pracovní výšky cca 25 cm.
7. Zmačknout start (počkat 10 min. . Může se pracovat
8. Clean box interior with tissue wetted in ethanol.
9. Turn the key to control red light on (This red light used to be on without interuption).
10. On the laptop press side button a keep for few second, until the green battery appears on screen.
11. Press the button again (to load software).
12. Press the arrow down and close the box completely.
13. Press UV. (UV will be switch off automatically after 20 minutes).
14. Press the arrow up to be glass window in working position (cca 25 cm).
15. Press start and wait for cca 10 minutes before you start work in the box.
16. After you finish work clean box interior with tissue wetted in ethanol.